## Tabelle1
Eye tracking specifications
Eye tracking technique | Video-based pupil and corneal reflection eye tracking with dark and bright pupil illumination modes. Two cameras capture stereo images of both eyes for robust accurate measurement of eye gaze and eye position in 3D space, as well as pupil diameter.
Eye tracking mode | Tobii Pro Spectrum has four different eye tracking modes for different groups of primate species: Human Great ape Monkey Small monkey
Sampling frequency | 60 Hz, 120 Hz, 150 Hz, 300 Hz, 600 Hz, or 1200 Hz (max. frequency depends on hardware version)
Precision¹ | Head supported² 0.01° RMS at optimal conditions (applying Savitzky-Golay filtering, settings listed in the test report) 0.06° RMS at optimal conditions (raw signal)
Head free-to-move³ 0.02° RMS at optimal conditions (applying Savitzky-Golay filtering, settings listed in the test report) 0.08° RMS at optimal conditions (raw signal)
Accuracy¹ | Head supported² Down to 0.16° Median: 0.30° (P25: 0.24° - P75: 0.41°)
Head free-to-move³ Down to 0.17° Median: 0.43° (P25: 0.32° - P75: 0.59°)
Binocular eye tracking | Yes
Eye tracker latency | Mean latency < 2 ms @ 1200Hz (SD < 0.2 ms)⁴
Blink recovery time | 1 frame (immediate)
Gaze recovery time | Less than 150 ms
Data sample output | Timestamp Gaze origin Gaze point Pupil diameter
Eye openness data stream | Eye openness stream has the same frequency as frequency selected for the gaze stream and will have the same timestamps. Eye openness data is provided in millimeters, for each eye individually.
Eye image data stream | Eye image stream frequency is approximately 10 Hz (one image with both eyes). Zoomed-in eye images available in tracking mode. Full-frame camera images are available in gaze recovery mode.
TTL input stream | 8–bit timestamped data (256 event codes) Event driven detection with a timestamp accuracy of 50 µs
Tracker and client time synchronization | Integrated between the eye tracker time domain and the client computer time domain with an accuracy of 100 µs
User calibration | Binocular, Monocular
Setup
Head movement tolerance | Excellent Dual-camera system, with more images than a one camera system, gives a more accurate data calculation and the best level of precision and robustness for head movement.
Operating distance (mounted on screen) | 55 to 75 cm (22" and 30") from the eye tracker
Freedom of head movement (width x height) | Freedom of head movement at 65cm distance 34 cm x 26 cm (13.5" x 10")
Freedom of head movement at 75cm distance 42 cm x 26 cm (16.5" x 10")
Optimal screen size | Up to 24" (16:9 aspect ratio)
Recommended monitor | Supplied 23.8" monitor
Tracker setup options | Eye tracker unit can be used alone or together with the monitor or other screens.
Can also be used with real-world stimuli.
Tracker mounted on stand, allows for even larger screens and physical objects to be tracked.
Participant setup | Very simple and easy. Typically, less than 1 minute.
Software and framework compatibility
Software and framework compatibility | Tobii Pro Lab Tobii Pro Eye Tracker Manager Tobii Pro SDK Any applications built on the Tobii Pro SDK
Operating system | Windows, macOS, Linux
Hardware versions
Hardware versions | 300 Hz 600 Hz 1200 Hz
Eye tracker unit
Dimensions (L x H x W) | 55 cm × 18 cm × 6 cm (22" × 7" × 2") The eye tracker is mounted on a stand, which raises it from the surface by 9 cm (approximately 4")
Weight | 5.1 kg (11.4 lbs.) With the power supply unit, the weight is 5.7 kg (12.9 lbs.)
Connectors | TTL input: 8-bit (DB-9 connector) Communication: Ethernet (RJ-45 connector) Power supply: 24 VDC (5.5 mm connector)
Data Processing | Fully embedded data processing
Eye tracking cameras | N/A
Illuminators | Dark pupil illumination modules Bright pupil illumination modules
Power consumption | Typical power consumption: 60 W Max. rated power consumption: 96 W
Power options | Input: 100-240 VAC 50/60 Hz Max. rated power consumption: 120 W No load power consumption: <0.15 W Energy efficiency level: VI Complies with EISA 2007/DoE, NRCan, AU/NZ MEPS, EU ErP, and CoC Version 5